- Enclosure No.: 1 to Specification of Essentials terms of the public contract (SIWZ)

Specification of the Subject

|  | Parameter | Required | Offered |
| --- | --- | --- | --- |
| 1 | Modulator type | „Solid-state” |  |
| 2 | Microwave power source | Klystron |  |
|  | Microwave power parameters |  |  |
| 3 | Frequency | Required range from 2997 to 3000 MHz |  |
| 4 | Peak power | Required range from od 2 to 7MW |  |
| 5 | Pulse length | Required range from 0,5 to 16us |  |
| 6 | Pulse repetition rate | Required range from 0,1 to 400 Hz |  |
| 7 | Average power | Required up to 36kW |  |
| 8 | Microwave power additional equipment | Solenoid, ionic pump, all required junction for connecting modulator with microwave source |  |
|  | Klystron modulator parameters |  |  |
| 9 | Electric peak power | 16,7 MW |  |
| 10 | Main frequency oscillator | Built-in, required power range from 0 – to 100W min |  |
| 11 | Electric average power | Minimum 98 kW |  |
| 12 | Cathode voltage range | Required range from 0 do 145kV |  |
| 13 | Cathode current range | Required range from 0 do 115A |  |
| 14 | Pulse repetition rate | Required range from 0 to 430 Hz |  |
| 15 | Pulse flatness | ±2% Pic. 1 |  |
| 16 | Pulse stability RMS | 50ppm |  |
| 17 | Rate of rise | >= 110kV/us |  |
| 18 | Rate of fall | >= 110kV/us |  |
| 19 | Trig delay | >= 1.3us |  |
| 20 | Klystron filament voltage | Required range from 0 to 8V DC |  |
| 21 | Klystron filament current | Required range from 0 to 33A DC |  |
| 22 | Interface communication | Modbus TCP |  |
| 23 | Additional requirement | The RF-units are designed to be driven in a pulse train mode (2-5x pulses, 1-3µs) (Pic. 2) |  |
| 24 | Working quality control | Digitizer for Modulator Diagnostics |  |
| 25 | RF working quality control | Digitizer for RF Diagnostics |  |
| 26 | Interface for main oscillator control | Internal via GUI and external via RS232 or Ethernet |  |
| 27 | SF6 pressure measurement | Built-in, manage via GUI |  |
| 28 | Additional protection circuit | 4 interlock inputs for external devices |  |
| 29 | Temperature measurement | At least 2 measurement: for tank oil and klystron collector |  |
| 30 | Maximum dimensions | 1800x2500x1982 (width x length x height) |  |
| 31 | Maximum weight | 2500 kg (excluding: oil, klystron, solenoid and radiation shields). |  |
| 32 | Miscellaneous | Factory acceptance test (FAT) certificate |  |
| 33 |  | Operation and service manual |  |
| 34 |  | Training for operation and maintenance at FAT |  |

Pic. 1 Pulse shape definition

Pic. 2 Pulse train specification