##### Enclosure No.: 2

##### to Specification of Essentials terms of the public contract (SIWZ)

Nr referencyjny: AZP.270.47.2019

....................................................................

....................................................................

(contractors’s name and seat)

……………………………………………….......................

(place of business)

...........................................................

(place and date)

OFFER FORM

In reply to contract notice for delivery of superconductive radio-frequency accelerating cryomodules  we submit this offer

- I offer to realize the contract for the total price (according to Enclosure B to Offer Form):

| OFFERED PRICE NET | .................................................................PLN/EUR (read: ......................................................................... ..................................................................................................) |
| --- | --- |
| VAT *) | ..................................................................PLN/EUR (read: ......................................................................... ..................................................................................................) |
| OFFERED PRICE TOTAL | ........................................................................PLN/EUR (read: ......................................................................... ..................................................................................................) |

### *) If  applicable - according to valid provisions.

- I state the contract complies with requirements identified in Enclosure nr 1 to Specification of Essentials terms of the public contract. I enclose Enclosure A to Offer Form for confirmation.

- I state that we were informed about the Specification of Essential terms of the public contract (including the draft of the essential contract’s provisions) we do not raise any objections, and we accept conditions of these documents.

- I offer to perform the contract within the time specified in Specification of Essentials terms of the public contract.

- I declare that offered warranty (guarantee)  period is ……..…… months (the minimum required warranty (guarantee) period  is 12 months).

- I state I paid a deposit in amount 100 000,00 PLN in form  …………………

bank account nr 58102011270000190200073015, IBAN PL58102011270000190200073015, SWFT BPKOPLPW

- I declare that I am / am not an active VAT taxpayer **).

- I state that offered price, mentioned in point 1 of this offer form, includes all contractor’s costs related with the contract to be incurred by the Orderer in case of selection of this offer.

- In case of awarding the public contract, I declare to sign the contract in the place and date specified by the Orderer and according to the essential contract’s  provisions in enclosure No. 4.

- We intend to execute the subject matter of the contract on our own/with the subcontractors **).  The list of some lots of the contract entrusted to subcontractors and the names of the subcontractors’ firms  - in the enclosure.

| Lp. | The name and the address of the subcontractor | VAT identification number/personal number National Court Register number/Central Registration and Information on Business number | The lot of the contract to be realized by the subcontractor |
| --- | --- | --- | --- |

- I state that I am bound by the offer for the period of 60 days starting from the deadline for submission of offers.

- Offer with enclosures is submitted on pages ………..

- I hereby declare, that information given on pages ......... are business secrets in terms of the act on elimination of unfair competition and cannot be published elsewhere.

/The Orderer’s requirement: The Contractor have to attach to the offer  the  factual and legal grounds that reserved information remains a business secret/.

- I declare, I am  microenterprise/small enterprise/medium-sized enterprise **) in the meaning of Act of 2 July 2004 on freedom of economic activity (the Journal of Law of 2017  item  2168 unified text with subsequent changes)

**) - delete as appropriate

- Teleaddress data for correspondence:

The name of the firm ***):

Address ***)……………..	...

tel. ***) 	 , fax. ***)

e-mail***) …………………………………………………………

VAT identification number:……………………………………………. National Official Business Register, …………………………………………

The data of the contact person:…………………………………………………………………………..

***) In case of the offer submitted jointly the dates of the appointed  plenipotentiary must be given

- I enclose the documents:

1) signed Enclosure No A to Specification of Essentials terms of the public contract (SIWZ)

2) ……………………………………………...............................................................

3) ……………………………………………...............................................................

- Enclosure No.: A

Specification of the Subject

| No. | Parameter | Required | Offered |
| --- | --- | --- | --- |
| General requirements |  |  |  |
| 1 | Amount | 3 pieces of Rossendorf type SRF Cryomodules |  |
| 2 | Delivery (Incoterms 2010) | DDP Narodowe Centrum Badań Jądrowych Andrzeja Sołtana 7 05-400 Otwock Poland |  |
| 3 | Protection during transportation | Cryomodules must be transported in protective frame preventing damage due to hard impacts |  |
| 4 | Installation | Performed jointly by the Contractor and the Orderer |  |
| 5 | Commissioning | Performed jointly by the Contractor and the Orderer |  |
| 6 | Warranty | Minimum 12 months from the date of SAT |  |
| Delivery dates |  |  |  |
|  | Overall delivery schedule | Partial delivery in accordance with overall PolFEL schedule. Each delivery date agreed with the Orderer. |  |
| Stage 1 |  |  |  |
|  | Delivery of a preliminary design documentation | Not later than 3 months after the contract date |  |

| Stage 2 |  |  |  |
| --- | --- | --- | --- |
|  | Component tests: Superconductive Cavities | Vertical tests: Not later than 18 months after the contract date |  |
| Stage 3 |  |  |  |
|  | Component tests: Fundamental Power Couplers | Warm tests: Not later than 21 months after the contract date |  |
| Stage 4 |  |  |  |
|  | Factory Acceptance Tests: Cryomodule 1 | Not later than 25 months after the contract date |  |
|  | Delivery of Cryomodule 1 | Not later than 27 months after the contract date |  |
| Stage 5 |  |  |  |
|  | Commissioning of Cryomodule 1 | Not later than 4 months after delivery |  |
| Stage 6 |  |  |  |
|  | Factory Acceptance Tests: Cryomodule 2 | Not later than 29 months after the contract date |  |
|  | Delivery of Cryomodule 2 | Not later than 31 months after the contract date |  |
| Stage 7 |  |  |  |
|  | Commissioning of Cryomodule 2 | Not later than 3 months after delivery |  |
| Stage 8 |  |  |  |
|  | Factory Acceptance Tests: Cryomodule 3 | Not later than 31 months after the contract date |  |
|  | Delivery of Cryomodule 3 | Not later than 33 months after the contract date |  |
| Stage 9 |  |  |  |
|  | Commissioning of Cryomodule 3 | Not later than 3 months after delivery |  |
|  | Final Site Acceptance Tests (SAT) | Not later than 36 months after the contract date |  |

| Design work |  |  |  |
| --- | --- | --- | --- |
| 1 | Redesign of the Cryomodule necessary to accommodate: | - Increased accelerating performance - Integration of modified TTF-III FPC - 5 K circuit for FPC intercept - Improved microphonic behavior (jointly with the Orderer) |  |
| 2 | Requirements in the area of microphonic behavior | 1. The Contractor shall prepare and make available to the Orderer full 3D CAD model of the cryomodule the design of which shall be aimed at keeping mechanical vibrations below 24Hz peak-to-peak 2. The Orderer shall have an option to verify, in agreement with the overall time schedule, the microphonic behavior of the model by means of FEA simulations and to propose eventual changes to the design. The changes must be realistic in terms of effort and costs and are subject to the approval by the Contractor. 3. The Contractor shall manufacture and assemble the cryomodules according to the modified or, if no changes were proposed by the Orderer, to the original design. |  |

| Detailed requirements |  |  |  |
| --- | --- | --- | --- |
| Cryomodule |  |  |  |
| 1 | Cryomodule type | Modified Rossendorf style, housing 2 Superconducting Radio Frequency 9-cell accelerating TESLA cavities, equipped with RF power couplers, HOM couplers and field probes |  |
| 2 | External dimension (L x W x H) | 3.5 x 1.1 x 2.6 m |  |
| 3 | Weight | Within the range: 1400 - 1500 kg |  |
| 4 | Cryomodule main components | 1. Cavity string with tuner systems 2. Helium system 3. Thermal shield with Nitrogen cooling system 4. Input couplers 5. Cryostat |  |
| 5 | Cryomodule external interfaces | 1. Electrical connection for tuner (2 pieces) 2. Electrical connection for sensors and heater 3. Coaxial connection for RF and HOM antennas 4. Cryogenic connection for LN2 supply 5. Cryogenic connection for N2 return 6. Connection to LN2 level sensor 7. Cryogenic connection for helium (2 K and possibly 5 K supply and return) 8. Vacuum vessel port for evacuation |  |

| 6 | Cryomodule external components | 1. Support frame 2. Transport rolls (4 pieces) 3. Module posts (4 pieces) 4. Horizontal adjustment mechanism (transversal/longitudinal) (4 x 2 pieces) 5. Mechanism for axial adjustment of cavity string (cold mass) (2 pieces) 6. Mechanism for radial adjustment of cavity string (cold mass) (6 pieces) |  |
| --- | --- | --- | --- |
| Cryostat |  |  |  |
| 7 | Vacuum vessel material | 304 stainless steel |  |
| 8 | Vacuum vessel volume | Within the range: 2150 - 2250 l |  |
| 9 | Cryomodule magnetic shield | Two layer: - outer, made of MuMetal or equivalent, operated at room temperature, - inner, made of Cryoperm or equivalent, operated at 2 K |  |
| 10 | Cryomodule magnetic shield performance | Damping factor > 100, Bres < 5 mG (at room temperatures) |  |
| Cavity string |  |  |  |
| 11 | Cavity operating frequency (at 2 K) | 1300.000 (+/- 0.05) MHz |  |
| 12 | Cavity and flanges material | Cavity - RRR300 niobium, Flanges - NbTi |  |
| 13 | Cavity preparation | Documented XFEL recipe (EP, 800 °C, 120 °C) |  |
| 14 | Cavity performance in vertical tests at 2 K | Accelerating gradient >23.5 MV/m, Quality factor Q0 >1.2 x 1010 |  |
| 15 | Helium vessel material | Titanium |  |
| 16 | Helium vessel volume | Volume within the range: 55 - 60 l |  |
| 17 | Helium vessel heater power | < 50 W per cavity |  |

| RF fundamental power coupler |  |  |  |
| --- | --- | --- | --- |
| 18 | Coupler type | Modified TTF-III style, produced according to XFEL specification with increased inner conductor copper plating thickness of 140 µm |  |
| 19 | Coupler external quality factor | Variable in the range 107 - 108 by means of a remotely controlled motorized tuner |  |
| 20 | Tuner design | XFEL design (motor and gear, piezo actuators) |  |
| 21 | Tuner range | +/- 150 kHz (by means of motor), 1 kHz (by means of piezo actuators) |  |
| 22 | Probe antenna external quality factor | Qext in the range 3 - 6 x 1011 |  |
| 23 | Coupler performance at warm tests | Documented minimum 8 hours stable operation (at room temperature) with 4 kW input coupler power to beam in the standing wave mode under the condition of continuous wave mode operation. The coupler vacuum must stay below 2 x 10-6 mbar. |  |
| HOM coupler |  |  |  |
| 24 | HOM-feedthroughs | XFEL-like |  |
| 25 | External quality factor | Qext > 1 x 1011 (pi-mode) |  |
| Cryomodule assembly |  |  |  |
| 26 | Assembly procedure requirements - cavities | Cavity assembly must be performed in ISO4 (or better) clean room |  |
| 27 | Assembly procedure requirements - cryomodule | Cryomodule assembly must be performed in ISO8 (or better) cleanroom |  |
| 28 | Magnetic field | Magnetic flux B during the cryomodule assembly cannot exceed 5mG (at cavity position prior to string integration) |  |

| Cryomodule performance during Factory Acceptance Test |  |  |  |
| --- | --- | --- | --- |
| Leak check at room temperature: |  |  |  |
| 29 | Cavity to ambient | < 2 x 10-10 mbar l/s |  |
| 30 | Cavity to helium vessel | < 2 x 10-10 mbar l/s |  |
| 31 | Helium vessel to insulation vacuum | < 2 x 10-8 mbar l/s |  |
| 32 | Insulation vacuum to ambient | < 1 x 10-7 mbar l/s |  |
| Coupler check |  |  |  |
| 33 | Coupler performance at warm tests | Minimum 8 hours stable operation (at room temperature) with 4 kW input coupler power to beam in the standing wave mode under the condition of continuous wave mode operation |  |
| 34 | Tuner test | Frequency change of: ± 20 kHz (by means of motor, in warm state) 1 kHz (by means of piezo actuators, in warm state) |  |
| Cryomodule performance during Site Acceptance Test |  |  |  |
| Leak check at room temperature: |  |  |  |
| 35 | Cavity to ambient | < 2 x 10-10 mbar l/s |  |
| 36 | Cavity to helium vessel | < 2 x 10-10 mbar l/s |  |
| 37 | Helium vessel to insulation vacuum | < 2 x 10-8 mbar l/s |  |
| 38 | Insulation vacuum to ambient | < 1 x 10-7 mbar l/s |  |
| Leak check at 2K (cooldown from RT to 4K/2K with NCBJ cryoplant and NCBJ control system) |  |  |  |
| 39 | Cavity to ambient | < 2 x 10-10 mbar l/s |  |
| 40 | Cavity to helium vessel | < 2 x 10-10 mbar l/s |  |
| 41 | Helium vessel to insulation vacuum | < 2 x 10-8 mbar l/s |  |
| 42 | Insulation vacuum to ambient | < 1 x 10-7 mbar l/s |  |

| Static losses at 2K |  |  |  |
| --- | --- | --- | --- |
| 43 | Static losses of each cryomodule | < 10 W in average (< 30 W for all 3 cryomodules) |  |
| Fundamental power coupler |  |  |  |
| 44 | Remotely controlled Qext | 1 x 107 to 1x 108 (tuner motor control system provided by the Orderer) |  |
| 45 | Stability of operation at 4 kW power (continuous wave) | Stable operation for at least 4 hours. The temperature rise induced by RF losses must be below 150 K. |  |
| 46 | Frequency | Tuner must achieve cold frequency of 1.3 GHz with tuning range at 1.3 GHz of ± 150 kHz |  |
| Cavity performance at 2K |  |  |  |
| 47 | Accelerating performance | Eacc > 18 MV/m (at Qext > 2,5 x 107) Q0 > 1.2 x 1010 in average for all 3 cryomodules (6 cavities), no individual cavity can have quality factor Q0 less than 1.0 x 1010 |  |
| 48 | Cryogenic load | The value of the cryogenic load directly corresponds to the quality factor. Total cryogenic load of all 3 cryomodules (6 cavities) must be below 200 W when operated at 18 MV/m and 2 K bath temperature. In addition no cavity should have a higher dynamic cryogenic load than 34 W ± 3 W under the same condition. |  |
| Positioning of cavities |  |  |  |
| 49 | Cavity positioning accuracy | Better than ± 300 µm, as demonstrated by means of laser tracker measurements |  |

| Microphonic behaviour |  |  |  |
| --- | --- | --- | --- |
| 50 | Mechanical vibration level of cryomodule | Mechanical vibrations must be below 24Hz peak-to-peak. Responsibility of the Contractor is limited to the cases that can be related to assembly mistakes. |  |
| Instrumentation of cryomodules |  |  |  |
| 51 | Status | Working |  |

Enclosure No.: B

PRICE FORM

| No | Stage | Maximal payment (percentage of the total net price) | Offered net price in PLN/ EUR |
| --- | --- | --- | --- |
| 1 | Advance | Up to 30% |  |
| 2 | Stage 1 - Delivery of a preliminary design documentation | Up to 20% |  |
| 3 | Stage 2 - Component tests: Superconductive Cavities | Up to 15% |  |
| 4 | Stage 3 - Component tests: Fundamental Power Couplers (FPC) | Up to 8% |  |
| 5 | Stage 4 - Factory Acceptance Tests (FAT) and Delivery of Cryomodule 1 | Up to 5 % |  |
| 6 | Stage 5 - Commissioning of Cryomodule 1 | Up to 4% |  |
| 7 | Stage 6 - Factory Acceptance Tests (FAT) and Delivery of Cryomodule 2 | Up to 5% |  |
| 8 | Stage 7 - Commissioning of Cryomodule 2 | Up to 4% |  |
| 9 | Stage 8 - Factory Acceptance Tests (FAT) and Delivery of Cryomodule 3 | Up to 5% |  |
| 10 | Stage 9 - Commissioning of Cryomodule 3, Final Site Acceptance Tests (SAT) |  |  |
| 11 | TOTAL |  |  |