- Enclosure No.: 1 to Specification of Essentials terms of the public contract (SIWZ)

Specification of the Subject

| O.n. | Parameter | Required | Offered |
| --- | --- | --- | --- |
| 1 | Modulator type | „Solid-state” |  |
| 2 | Microwave power source | Klystron |  |
| Microwave power source parameters |  |  |  |
| 3 | Frequency | Required range from 2997 to 3000 MHz |  |
| 4 | Peak power | Required range from od 2 to 7.5MW |  |
| 5 | Pulse length | Required range from 0,5 to 5.0µs |  |
| 6 | Average power | Required up to 5,6kW |  |
| 7 | Microwave power source additional equipment | Solenoid, ionic pump, all required junction for connecting modulator with microwave source |  |
| Klystron modulator parameters |  |  |  |
| 8 | Electric peak power | 18,6 MW |  |
| 9 | Main frequency oscillator | Built-in, required power range from 0 – to 100W min |  |
| 10 | Electric average power | Minimum 18 kW |  |
| 11 | Pulse repetition rate | Required range from 0,1 to 150 Hz - subject to max. average power, with the option of extending up to 300Hz after purchasing the appropriate equipment |  |
| 12 | Cathode voltage range | Required range from 0 do 160kV |  |
| 13 | Cathode current range | Required range from 0 do 115A |  |
| 14 | Pulse flatness | ±1,5% Pic. 1 |  |
| 15 | Pulse stability RMS | 100ppm |  |
| 16 | Rate of rise | 140kV/µs±5% |  |
| 17 | Rate of fall | 140kV/µs±5% |  |
| 18 | Trig delay | 1.2µs±10% |  |
| 19 | Klystron filament voltage | Required range from 0 to 20V DC |  |
| 20 | Klystron filament current | Required range from 0 to 18A DC |  |
| 21 | Interface communication | Modbus TCP |  |
| 22 | Working quality control | Digitizer for Modulator Diagnostics |  |
| 23 | Interface for main oscillator control | Internal via GUI and external via RS232 or Ethernet |  |
| 24 | Additional protection circuit | 4 interlock inputs for external devices |  |
| 25 | Temperature measurement | At least 2 measurement: for tank oil and klystron collector |  |
| 26 | Maximum dimensions | 1600x650x2100 (width x length x height) |  |
| 27 | Maximum weight | 1800 kg (excluding: oil, klystron, solenoid and radiation shields). |  |
| 28 | Miscellaneous | Factory acceptance test (FAT) certificate |  |
|  |  | Operation and service manual |  |
|  |  | Training for operation and maintenance at FAT (online training is allowed) |  |

Pic. 1 Pulse shape definition